## Raw data
May 8 - 16 flow/ pressure measurements | [29 empty] | b
Meter: FL3440ST: 41512 sccm at 150 mm |  |  |  | Blank rig |  |  |  | H400, double-ply (white and blue) | [6 empty] | H400, quadruple-ply | [6 empty] | H400 (single-ply, blue only) + HVMerv14 |  |  |  |  | H400 (double-ply) + HVMerv14 | [5 empty] | H300 (single-ply, blue only) + HVMerv14 |  |  |  |  | H300 (double ply) +HV Merv 14 | [5 empty] | H600 (single-ply, blue only) + HVMerv14 |  |  |  |  | H600 (double-ply) + HVMerv 14 | [5 empty] | H600, (quadruple-ply) | [5 empty] | HVMerv 14, double-ply |  |  |  |  | H100 (single-ply, blue only) + HVMerv14 |  |  |  |  | H100 (double ply) + HVMerv 14 |  |  |  |  | H100, (quadruple-ply) |  |  |  |  | Camfil HEPA filter |  |  |  |  | Camfil HEPA, 2-ply |  |  |  |  | Woodbridge 5-ply |  |  |  |  | Q1 |  |  |  |  | Q2 |  |  |  |  | Q3 |  |  |  |  | Q4 |  |  |  |  | Q5 |  |  |  |  | Yellow pad, 2-layer (yellow-white-white-yellow) |  |  |  |  | Yellow pad, 4-layer (4 x yellow-white back to back) |  |  |  |  | Yellow pad, 1-layer (yellow-white flow from yellow to white) |  |  |  |  | Retest of blank rig, new seals |  |  | Yellow pad, 2-layer (retest) |  |  |  |  | Q3 (retest)
2020-05-16
Meter | SCCM | SLPM | SLPM/M2 | h_l | h_r | in_water | previous | h_l | h_r | in_water | corrected in_water | previous rig | mm H20 |  | h_l | h_r | in_water | corrected in_water | previous rig | mm H20 |  | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | previous rig | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | previous rig | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | previous rig | mm H20 | h_l | h_r | in_water | corrected in_water | previous rig | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | h_l | h_r | in_water | corrected in_water | mm H20 | h_l | h_r | in_water | corrected in_water | mm H20
0 | 0 | 0 | 0 | 44.6 | 55.4 | 0 | 0 | 44.6 | 55.3 | 0 | 0 | 0 | 0 |  | 44.6 | 55.3 | 0 | 0 | 0 | 0 |  | 47.15 | 52.7 | 0 | 0 | 0 | 44.6 | 55.3 | 0 | 0 | 0 | 0 | 47.15 | 52.7 | 0 | 0 | 0 | 44.625 | 55.3 | 0 | 0 | 0 | 0 | 47.2 | 52.7 | 0 | 0 | 0 | 44.675 | 55.3 | 0 | 0 | 0 | 0 | 44.7 | 55.3 | 0 | 0 | 0 | 0 | 45.45 | 54.35 | 0 | 0 | 0 | 47.2 | 52.7 | 0 | 0 | 0 | 45.5 | 54.35 | 0 | 0 | 0 | 45.5 | 54.35 | 0 | 0 | 0 | 45.55 | 54.35 | 0 | 0 | 0 | 45.55 | 54.35 | 0 | 0 | 0 | 45.55 | 54.35 | 0 | 0 | 0 | 44.5 | 55.45 | 0 | 0 | 0 | 44.5 | 55.45 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 46.15 | 53.7 | 0 | 46.2 | 53.55 | 0 | 0 | 0 | 46.15 | 53.7 | 0 | 0 | 0
5 | 1383.7333 | 1.3837 | 1100.8221 | 44.6 | 55.4 | 0 | 0 | 44.6 | 55.3 | 0 | 0 | 0 | 0 |  | 44.65 | 55.35 | 0.03937 | 0.03937 | 0.1 | 1 |  | 47.15 | 52.75 | 0.019685 | 0.019685 | 0.5 | 44.65 | 55.35 | 0.03937 | 0.03937 | 0.2 | 1 | 47.15 | 52.7 | 0 | 0 | 0 | 44.65 | 55.3 | 0.0098425 | 0.0098425 | 0.1 | 0.25 | 47.2 | 52.65 | -0.019685 | -0.019685 | -0.5 | 44.7 | 55.3 | 0.0098425 | 0.0098425 | 0.1 | 0.25 | 44.75 | 55.3 | 0.019685 | 0.019685 | 0.1 | 0.5 | 45.45 | 54.35 | 0 | 0 | 0 | 47.2 | 52.7 | 0 | 0 | 0 | 45.5 | 54.4 | 0.019685 | 0.019685 | 0.5 | 45.5 | 54.35 | 0 | 0 | 0 | 45.7 | 54.55 | 0.1378 | 0.1378 | 3.5 | 45.9 | 54.7 | 0.27559 | 0.27559 | 7 | 45.6 | 54.4 | 0.03937 | 0.03937 | 1 | 44.5 | 55.45 | 0 | 0 | 0 | 44.5 | 55.45 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.5 | 55.4 | 0 | 0 | 0 | 44.55 | 55.4 | 0.019685 | 0.019685 | 0.5 | 44.5 | 55.4 | 0 | 0 | 0 | 46.15 | 53.7 | 0 | 46.3 | 53.6 | 0.059055 | 0.059055 | 1.5 | 46.15 | 53.7 | 0 | 0 | 0
10 | 2767.4667 | 2.7675 | 2201.6441 | 44.6 | 55.4 | 0 | 0 | 44.75 | 55.5 | 0.1378 | 0.1378 | 0.2 | 3.5 |  | 44.95 | 55.6 | 0.25591 | 0.25591 | 0.4 | 6.5 |  | 47.25 | 52.9 | 0.11811 | 0.11811 | 3 | 44.9 | 55.55 | 0.21654 | 0.21654 | 0.3 | 5.5 | 47.3 | 52.85 | 0.11811 | 0.11811 | 3 | 44.85 | 55.5 | 0.16732 | 0.16732 | 0.2 | 4.25 | 47.3 | 52.85 | 0.098425 | 0.098425 | 2.5 | 44.95 | 55.5 | 0.18701 | 0.18701 | 0.2 | 4.75 | 45 | 55.6 | 0.23622 | 0.23622 | 0.3 | 6 | 45.7 | 54.6 | 0.19685 | 0.19685 | 5 | 47.3 | 52.85 | 0.098425 | 0.098425 | 2.5 | 45.7 | 54.55 | 0.15748 | 0.15748 | 4 | 45.7 | 54.55 | 0.15748 | 0.15748 | 4 | 46.2 | 55 | 0.51181 | 0.51181 | 13 | 46.8 | 55.55 | 0.96457 | 0.96457 | 24.5 | 45.9 | 54.7 | 0.27559 | 0.27559 | 7 | 44.6 | 55.55 | 0.07874 | 0.07874 | 2 | 44.5 | 55.45 | 0 | 0 | 0 | 44.55 | 55.5 | 0.059055 | 0.059055 | 1.5 | 44.5 | 55.4 | 0 | 0 | 0 | 44.6 | 55.55 | 0.098425 | 0.098425 | 2.5 | 44.6 | 55.45 | 0.059055 | 0.059055 | 1.5 | 44.8 | 55.65 | 0.21654 | 0.21654 | 5.5 | 44.6 | 55.4 | 0.03937 | 0.03937 | 1 | 46.15 | 53.7 | 0 | 46.35 | 53.6 | 0.07874 | 0.07874 | 2 | 46.25 | 53.75 | 0.059055 | 0.059055 | 1.5
15 | 4151.2 | 4.1512 | 3302.4662 | 44.6 | 55.4 | 0 | 0.1 | 44.9 | 55.6 | 0.23622 | 0.23622 | 0.3 | 6 |  | 45.25 | 55.9 | 0.49213 | 0.49213 | 0.6 | 12.5 |  | 47.5 | 53.05 | 0.27559 | 0.27559 | 7 | 45.15 | 55.8 | 0.41339 | 0.41339 | 0.5 | 10.5 | 47.45 | 53 | 0.23622 | 0.23622 | 6 | 45.05 | 55.65 | 0.30512 | 0.30512 | 0.4 | 7.75 | 47.5 | 53 | 0.23622 | 0.23622 | 6 | 45.15 | 55.7 | 0.34449 | 0.34449 | 0.4 | 8.75 | 45.3 | 55.8 | 0.43307 | 0.43307 | 0.5 | 11 | 46 | 54.85 | 0.41339 | 0.41339 | 10.5 | 47.5 | 53 | 0.23622 | 0.23622 | 6 | 46.1 | 54.75 | 0.3937 | 0.3937 | 10 | 45.95 | 54.8 | 0.35433 | 0.35433 | 9 | 46.7 | 55.45 | 0.88583 | 0.88583 | 22.5 | 47.75 | 56.45 | 1.6929 | 1.6929 | 43 | 46.25 | 55 | 0.5315 | 0.5315 | 13.5 | 44.75 | 55.7 | 0.19685 | 0.19685 | 5 | 44.55 | 55.5 | 0.03937 | 0.03937 | 1 | 44.7 | 55.65 | 0.17717 | 0.17717 | 4.5 | 44.6 | 55.5 | 0.07874 | 0.07874 | 2 | 44.75 | 55.7 | 0.21654 | 0.21654 | 5.5 | 44.7 | 55.6 | 0.15748 | 0.15748 | 4 | 45 | 55.85 | 0.37402 | 0.37402 | 9.5 | 44.6 | 55.45 | 0.059055 | 0.059055 | 1.5 | 46.15 | 53.7 | 0 | 46.5 | 53.75 | 0.19685 | 0.19685 | 5 | 46.4 | 53.9 | 0.17717 | 0.17717 | 4.5
20 | 5534.9333 | 5.5349 | 4403.2883 | 44.6 | 55.4 | 0 | 0.2 | 45.05 | 55.75 | 0.35433 | 0.35433 | 0.4 | 9 |  | 45.45 | 56.15 | 0.66929 | 0.66929 | 0.8 | 17 |  | 47.75 | 53.25 | 0.45276 | 0.45276 | 11.5 | 45.35 | 56 | 0.57087 | 0.57087 | 0.6 | 14.5 | 47.6 | 53.15 | 0.35433 | 0.35433 | 9 | 45.25 | 55.85 | 0.4626 | 0.4626 | 0.5 | 11.75 | 47.65 | 53.2 | 0.37402 | 0.37402 | 9.5 | 45.35 | 55.95 | 0.52165 | 0.52165 | 0.6 | 13.25 | 45.5 | 56.1 | 0.62992 | 0.62992 | 0.7 | 16 | 46.15 | 55 | 0.5315 | 0.5315 | 13.5 | 47.6 | 53.15 | 0.33465 | 0.33465 | 8.5 | 46.15 | 54.95 | 0.49213 | 0.49213 | 12.5 | 46.2 | 55 | 0.5315 | 0.5315 | 13.5 | 47.1 | 55.9 | 1.2205 | 1.2205 | 31 | 48.7 | 57.35 | 2.4213 | 2.4213 | 61.5 | 46.6 | 55.35 | 0.80709 | 0.80709 | 20.5 | 44.9 | 55.8 | 0.29528 | 0.29528 | 7.5 | 44.6 | 55.55 | 0.07874 | 0.07874 | 2 | 44.8 | 55.75 | 0.25591 | 0.25591 | 6.5 | 44.7 | 55.65 | 0.17717 | 0.17717 | 4.5 | 44.9 | 55.85 | 0.33465 | 0.33465 | 8.5 | 44.85 | 55.7 | 0.25591 | 0.25591 | 6.5 | 45.25 | 56.1 | 0.57087 | 0.57087 | 14.5 | 44.7 | 55.55 | 0.1378 | 0.1378 | 3.5 | 46.15 | 53.7 | 0 | 46.6 | 53.85 | 0.27559 | 0.27559 | 7 | 46.5 | 54 | 0.25591 | 0.25591 | 6.5
25 | 6918.6667 | 6.9187 | 5504.1103 | 44.65 | 55.4 | 0.019685 | 0.3 | 45.2 | 55.95 | 0.49213 | 0.47244 | 0.5 | 12 |  | 45.75 | 56.45 | 0.90551 | 0.88583 | 0.9 | 22.5 |  | 47.95 | 53.5 | 0.62992 | 0.61024 | 15.5 | 45.6 | 56.25 | 0.76772 | 0.74803 | 0.8 | 19 | 47.75 | 53.5 | 0.55118 | 0.5315 | 13.5 | 45.6 | 56.05 | 0.67913 | 0.65945 | 0.6 | 16.75 | 47.85 | 53.4 | 0.5315 | 0.51181 | 13 | 45.6 | 56.2 | 0.7185 | 0.69882 | 0.7 | 17.75 | 45.8 | 56.35 | 0.84646 | 0.82677 | 0.8 | 21 | 46.4 | 55.25 | 0.72835 | 0.70866 | 18 | 47.8 | 53.3 | 0.47244 | 0.45276 | 11.5 | 46.3 | 55.15 | 0.62992 | 0.61024 | 15.5 | 46.45 | 55.2 | 0.70866 | 0.68898 | 17.5 | 47.6 | 56.4 | 1.6142 | 1.5945 | 40.5 | 49.6 | 58.3 | 3.1496 | 3.1299 | 79.5 | 46.9 | 55.65 | 1.0433 | 1.0236 | 26 | 45.05 | 55.95 | 0.41339 | 0.3937 | 10 | 44.65 | 55.65 | 0.1378 | 0.11811 | 3 | 44.95 | 55.85 | 0.35433 | 0.33465 | 8.5 | 44.75 | 55.7 | 0.21654 | 0.19685 | 5 | 45.05 | 55.95 | 0.43307 | 0.41339 | 10.5 | 44.95 | 55.85 | 0.35433 | 0.33465 | 8.5 | 45.5 | 56.3 | 0.74803 | 0.72835 | 18.5 | 44.75 | 55.65 | 0.19685 | 0.17717 | 4.5 | 46.15 | 53.7 | 0 | 46.8 | 54 | 0.41339 | 0.41339 | 10.5 | 46.65 | 54.15 | 0.37402 | 0.37402 | 9.5
30 | 8302.4 | 8.3024 | 6604.9324 | 44.7 | 55.4 | 0.03937 | 0.5 | 45.6 | 56.1 | 0.70866 | 0.66929 | 0.6 | 17 |  | 46 | 56.7 | 1.1024 | 1.063 | 1.2 | 27 |  | 48.15 | 53.7 | 0.7874 | 0.74803 | 19 | 45.85 | 56.5 | 0.96457 | 0.9252 | 1 | 23.5 | 47.95 | 53.45 | 0.61024 | 0.57087 | 14.5 | 45.6 | 56.25 | 0.75787 | 0.7185 | 0.7 | 18.25 | 48.05 | 53.6 | 0.68898 | 0.64961 | 16.5 | 45.85 | 56.5 | 0.93504 | 0.89567 | 0.8 | 22.75 | 46.1 | 56.65 | 1.0827 | 1.0433 | 1 | 26.5 | 46.7 | 55.5 | 0.94488 | 0.90551 | 23 | 47.95 | 53.5 | 0.61024 | 0.57087 | 14.5 | 46.55 | 55.35 | 0.80709 | 0.76772 | 19.5 | 46.6 | 55.45 | 0.86614 | 0.82677 | 21 | 48.15 | 56.9 | 2.0276 | 1.9882 | 50.5 | 50.5 | 59.2 | 3.8583 | 3.8189 | 97 | 47.3 | 56.05 | 1.3583 | 1.3189 | 33.5 | 45.2 | 56.1 | 0.5315 | 0.49213 | 12.5 | 44.7 | 55.7 | 0.17717 | 0.1378 | 3.5 | 45.1 | 56 | 0.47244 | 0.43307 | 11 | 44.85 | 55.8 | 0.29528 | 0.25591 | 6.5 | 45.2 | 56.15 | 0.57087 | 0.5315 | 13.5 | 45.1 | 55.95 | 0.45276 | 0.41339 | 10.5 | 45.7 | 56.55 | 0.9252 | 0.88583 | 22.5 | 44.85 | 55.7 | 0.25591 | 0.21654 | 5.5 | 46.2 | 53.75 | 0.03937 | 46.9 | 54.15 | 0.51181 | 0.47244 | 12 | 46.8 | 54.25 | 0.47244 | 0.43307 | 11
35 | 9686.1333 | 9.6861 | 7705.7544 | 44.75 | 55.45 | 0.07874 | 0.6 | 45.6 | 56.25 | 0.76772 | 0.68898 | 0.7 | 17.5 |  | 46.4 | 57.05 | 1.3976 | 1.3189 | 1.4 | 33.5 |  | 48.4 | 53.85 | 0.94488 | 0.86614 | 22 | 46.1 | 56.8 | 1.1811 | 1.1024 | 1.3 | 28 | 48.15 | 53.65 | 0.76772 | 0.68898 | 17.5 | 45.85 | 56.45 | 0.93504 | 0.8563 | 0.9 | 21.75 | 48.25 | 53.75 | 0.82677 | 0.74803 | 19 | 46.1 | 56.7 | 1.1122 | 1.0335 | 1.1 | 26.25 | 46.4 | 56.9 | 1.2992 | 1.2205 | 1.4 | 31 | 46.9 | 55.75 | 1.122 | 1.0433 | 26.5 | 48.15 | 53.65 | 0.74803 | 0.66929 | 17 | 46.8 | 55.6 | 1.0039 | 0.9252 | 23.5 | 46.9 | 55.7 | 1.0827 | 1.0039 | 25.5 | 48.6 | 57.35 | 2.3819 | 2.3031 | 58.5 | 51.45 | 60.1 | 4.5866 | 4.5079 | 114.5 | 47.6 | 56.35 | 1.5945 | 1.5157 | 38.5 | 45.4 | 56.3 | 0.68898 | 0.61024 | 15.5 | 44.8 | 55.75 | 0.23622 | 0.15748 | 4 | 45.2 | 56.15 | 0.57087 | 0.49213 | 12.5 | 45 | 55.9 | 0.3937 | 0.31496 | 8 | 45.4 | 56.3 | 0.70866 | 0.62992 | 16 | 45.25 | 56.1 | 0.57087 | 0.49213 | 12.5 | 46 | 56.8 | 1.1417 | 1.063 | 27 | 44.95 | 55.75 | 0.31496 | 0.23622 | 6 | 46.25 | 53.8 | 0.07874 | 47.05 | 54.3 | 0.62992 | 0.55118 | 14 | 46.9 | 54.4 | 0.57087 | 0.49213 | 12.5
40 | 11069.8667 | 11.0699 | 8806.5765 | 44.8 | 55.5 | 0.11811 | 0.8 | 45.7 | 56.4 | 0.86614 | 0.74803 | 0.8 | 19 |  | 46.7 | 57.35 | 1.6339 | 1.5157 | 1.6 | 38.5 |  | 48.65 | 54.15 | 1.1614 | 1.0433 | 26.5 | 46.45 | 57.05 | 1.4173 | 1.2992 | 1.4 | 33 | 48.35 | 53.85 | 0.9252 | 0.80709 | 20.5 | 46.05 | 56.7 | 1.1122 | 0.99409 | 1 | 25.25 | 48.5 | 54.05 | 1.0433 | 0.9252 | 23.5 | 46.45 | 56.95 | 1.3484 | 1.2303 | 1.2 | 31.25 | 46.75 | 57.25 | 1.5748 | 1.4567 | 1.4 | 37 | 47.2 | 56 | 1.3386 | 1.2205 | 31 | 48.4 | 53.85 | 0.9252 | 0.80709 | 20.5 | 47 | 55.85 | 1.1811 | 1.063 | 27 | 47.15 | 55.95 | 1.2795 | 1.1614 | 29.5 | 49.1 | 57.8 | 2.7559 | 2.6378 | 67 | 52.5 | 61.4 | 5.5118 | 5.3937 | 137 | 48 | 56.7 | 1.8898 | 1.7717 | 45 | 45.5 | 56.45 | 0.7874 | 0.66929 | 17 | 44.9 | 55.85 | 0.31496 | 0.19685 | 5 | 45.4 | 56.3 | 0.70866 | 0.59055 | 15 | 45.1 | 56.05 | 0.49213 | 0.37402 | 9.5 | 45.55 | 56.45 | 0.82677 | 0.70866 | 18 | 45.4 | 56.25 | 0.68898 | 0.57087 | 14.5 | 46.3 | 57.1 | 1.378 | 1.2598 | 32 | 45.05 | 55.9 | 0.41339 | 0.29528 | 7.5 | 46.3 | 53.85 | 0.11811 | 47.15 | 54.4 | 0.70866 | 0.59055 | 15 | 47.05 | 54.55 | 0.68898 | 0.57087 | 14.5
45 | 12453.6 | 12.4536 | 9907.3986 | 44.9 | 55.55 | 0.17717 | 1 | 45.9 | 56.65 | 1.0433 | 0.86614 | 0.9 | 22 |  | 47.05 | 57.75 | 1.9291 | 1.752 | 1.9 | 44.5 |  | 48.9 | 54.35 | 1.3386 | 1.1614 | 29.5 | 46.7 | 57.4 | 1.6535 | 1.4764 | 1.6 | 37.5 | 48.6 | 54.05 | 1.1024 | 0.9252 | 23.5 | 46.35 | 56.95 | 1.3287 | 1.1516 | 1.2 | 29.25 | 48.7 | 54.25 | 1.2008 | 1.0236 | 26 | 46.7 | 57.25 | 1.565 | 1.3878 | 1.4 | 35.25 | 47.05 | 57.6 | 1.8307 | 1.6535 | 1.7 | 42 | 47.5 | 56.3 | 1.5748 | 1.3976 | 35.5 | 48.6 | 54.05 | 1.0827 | 0.90551 | 23 | 47.25 | 56.05 | 1.3583 | 1.1811 | 30 | 47.4 | 56.2 | 1.4764 | 1.2992 | 33 | 49.65 | 58.35 | 3.189 | 3.0118 | 76.5 | 53.5 | 62.1 | 6.1811 | 6.0039 | 152.5 | 48.4 | 57.05 | 2.185 | 2.0079 | 51 | 45.7 | 56.65 | 0.94488 | 0.76772 | 19.5 | 45 | 55.95 | 0.3937 | 0.21654 | 5.5 | 45.65 | 55.55 | 0.51181 | 0.33465 | 8.5 | 45.25 | 56.15 | 0.59055 | 0.41339 | 10.5 | 45.75 | 56.65 | 0.98425 | 0.80709 | 20.5 | 45.5 | 56.4 | 0.7874 | 0.61024 | 15.5 | 46.55 | 57.35 | 1.5748 | 1.3976 | 35.5 | 45.15 | 56 | 0.49213 | 0.31496 | 8 | 46.4 | 53.9 | 0.17717 | 47.4 | 54.6 | 0.88583 | 0.70866 | 18 | 47.2 | 54.75 | 0.82677 | 0.64961 | 16.5
50 | 13837.3333 | 13.8373 | 11008.2206 | 45 | 55.65 | 0.25591 | 1.2 | 46.1 | 56.85 | 1.2008 | 0.94488 | 1 | 24 |  | 47.35 | 58.05 | 2.1654 | 1.9094 | 2.1 | 48.5 |  | 49.15 | 54.6 | 1.5354 | 1.2795 | 32.5 | 47.05 | 57.7 | 1.9094 | 1.6535 | 1.8 | 42 | 48.75 | 54.25 | 1.2402 | 0.98425 | 25 | 46.55 | 57.2 | 1.5059 | 1.25 | 1.4 | 31.75 | 48.95 | 54.45 | 1.378 | 1.122 | 28.5 | 47 | 57.55 | 1.8012 | 1.5453 | 1.7 | 39.25 | 47.4 | 58 | 2.126 | 1.8701 | 2 | 47.5 | 47.75 | 56.5 | 1.752 | 1.4961 | 38 | 48.8 | 54.3 | 1.2598 | 1.0039 | 25.5 | 47.5 | 56.3 | 1.5551 | 1.2992 | 33 | 47.75 | 56.45 | 1.7126 | 1.4567 | 37 | 50.15 | 58.9 | 3.6024 | 3.3465 | 85 | 54.55 | 63.25 | 7.0472 | 6.7913 | 172.5 | 48.7 | 57.45 | 2.4606 | 2.2047 | 56 | 45.85 | 56.8 | 1.063 | 0.80709 | 20.5 | 45.1 | 56.1 | 0.49213 | 0.23622 | 6 | 56.8 | 55.7 | 4.9606 | 4.7047 | 119.5 | 45.35 | 56.3 | 0.68898 | 0.43307 | 11 | 45.95 | 56.85 | 1.1417 | 0.88583 | 22.5 | 45.7 | 56.55 | 0.9252 | 0.66929 | 17 | 46.85 | 57.65 | 1.811 | 1.5551 | 39.5 | 45.25 | 56.1 | 0.57087 | 0.31496 | 8 | 46.45 | 54 | 0.23622 | 47.5 | 54.75 | 0.98425 | 0.74803 | 19 | 47.4 | 54.9 | 0.96457 | 0.72835 | 18.5
55 | 15221.0667 | 15.2211 | 12109.0427 | 45.05 | 55.7 | 0.29528 | 1.5 | 46.35 | 57.05 | 1.378 | 1.0827 | 1.1 | 27.5 |  | 47.7 | 58.45 | 2.4606 | 2.1654 | 2.3 | 55 |  | 49.55 | 55 | 1.8504 | 1.5551 | 39.5 | 47.45 | 58.05 | 2.2047 | 1.9094 | 2 | 48.5 | 49 | 54.45 | 1.4173 | 1.122 | 28.5 | 46.9 | 57.5 | 1.7618 | 1.4665 | 1.5 | 37.25 | 49.25 | 54.7 | 1.5945 | 1.2992 | 33 | 47.3 | 57.8 | 2.0177 | 1.7224 | 1.9 | 43.75 | 47.8 | 58.35 | 2.4213 | 2.126 | 2.2 | 54 | 48.15 | 56.9 | 2.0669 | 1.7717 | 45 | 49.05 | 54.5 | 1.437 | 1.1417 | 29 | 47.9 | 56.65 | 1.8504 | 1.5551 | 39.5 | 48.05 | 56.75 | 1.9488 | 1.6535 | 42 | 50.85 | 59.55 | 4.1339 | 3.8386 | 97.5 | 55.5 | 64 | 7.7165 | 7.4213 | 188.5 | 49.15 | 57.85 | 2.7953 | 2.5 | 63.5 | 46.1 | 56.95 | 1.2205 | 0.9252 | 23.5 | 45.2 | 56.15 | 0.55118 | 0.25591 | 6.5 | 45.9 | 55.8 | 0.70866 | 0.41339 | 10.5 | 45.5 | 56.4 | 0.7874 | 0.49213 | 12.5 | 46.15 | 57.05 | 1.2992 | 1.0039 | 25.5 | 45.9 | 56.7 | 1.063 | 0.76772 | 19.5 | 47.15 | 58 | 2.0669 | 1.7717 | 45 | 45.4 | 56.2 | 0.66929 | 0.37402 | 9.5 | 46.5 | 54.05 | 0.27559 | 47.7 | 54.95 | 1.1417 | 0.86614 | 22 | 47.55 | 55.05 | 1.0827 | 0.80709 | 20.5
60 | 16604.8 | 16.6048 | 13209.8648 | 45.1 | 55.8 | 0.35433 | 1.8 | 46.6 | 57.3 | 1.5748 | 1.2205 | 1.2 | 31 |  | 48.2 | 58.85 | 2.815 | 2.4606 | 2.5 | 62.5 |  | 49.7 | 55.2 | 1.9882 | 1.6339 | 41.5 | 47.8 | 58.4 | 2.4803 | 2.126 | 2.2 | 54 | 49.25 | 54.75 | 1.6339 | 1.2795 | 32.5 | 47.2 | 57.75 | 1.9783 | 1.624 | 1.6 | 41.25 | 49.5 | 54.95 | 1.7913 | 1.437 | 36.5 | 47.7 | 58.25 | 2.3524 | 1.998 | 1.9 | 50.75 | 48.2 | 58.75 | 2.7362 | 2.3819 | 2.3 | 60.5 | 48.45 | 57.2 | 2.3031 | 1.9488 | 49.5 | 49.3 | 54.7 | 1.6142 | 1.2598 | 32 | 48.2 | 56.9 | 2.0669 | 1.7126 | 43.5 | 48.35 | 57.05 | 2.185 | 1.8307 | 46.5 | 51.5 | 60.15 | 4.626 | 4.2717 | 108.5 | 56.8 | 65.4 | 8.7795 | 8.4252 | 214 | 49.65 | 58.35 | 3.189 | 2.8346 | 72 | 46.35 | 57.25 | 1.437 | 1.0827 | 27.5 | 45.35 | 56.3 | 0.66929 | 0.31496 | 8 | 46.1 | 57.05 | 1.2795 | 0.9252 | 23.5 | 45.7 | 56.55 | 0.9252 | 0.57087 | 14.5 | 46.4 | 57.25 | 1.4764 | 1.122 | 28.5 | 46.1 | 56.95 | 1.2402 | 0.88583 | 22.5 | 47.5 | 58.35 | 2.3425 | 1.9882 | 50.5 | 45.5 | 56.35 | 0.76772 | 0.41339 | 10.5 | 46.6 | 54.15 | 0.35433 | 47.9 | 55.15 | 1.2992 | 0.94488 | 24 | 47.8 | 55.25 | 1.2598 | 0.90551 | 23
65 | 17988.5333 | 17.9885 | 14310.6868 | 45.2 | 55.9 | 0.43307 | 2.2 | 46.85 | 57.55 | 1.7717 | 1.3386 | 1.3 | 34 |  | 48.55 | 59.25 | 3.1102 | 2.6772 | 2.6 | 68 |  | 50 | 55.5 | 2.2244 | 1.7913 | 45.5 | 48.15 | 58.75 | 2.7559 | 2.3228 | 2.2 | 59 | 49.5 | 54.9 | 1.7913 | 1.3583 | 34.5 | 47.5 | 58.1 | 2.2343 | 1.8012 | 1.7 | 45.75 | 49.7 | 55.25 | 1.9882 | 1.5551 | 39.5 | 48.05 | 58.55 | 2.6083 | 2.1752 | 2 | 55.25 | 48.55 | 59.05 | 2.9921 | 2.5591 | 2.4 | 65 | 48.85 | 57.55 | 2.5984 | 2.1654 | 55 | 49.5 | 54.95 | 1.7913 | 1.3583 | 34.5 | 48.5 | 57.25 | 2.3228 | 1.8898 | 48 | 48.6 | 57.4 | 2.4213 | 1.9882 | 50.5 | 52.1 | 60.8 | 5.1181 | 4.685 | 119 | 58 | 66.6 | 9.7244 | 9.2913 | 236 | 50.1 | 58.75 | 3.5236 | 3.0906 | 78.5 | 46.55 | 57.45 | 1.5945 | 1.1614 | 29.5 | 45.45 | 56.4 | 0.74803 | 0.31496 | 8 | 46.3 | 57.25 | 1.437 | 1.0039 | 25.5 | 45.85 | 56.7 | 1.0433 | 0.61024 | 15.5 | 46.6 | 57.5 | 1.6535 | 1.2205 | 31 | 46.25 | 57.1 | 1.3583 | 0.9252 | 23.5 | 47.85 | 58.7 | 2.6181 | 2.185 | 55.5 | 45.65 | 56.45 | 0.86614 | 0.43307 | 11 | 46.7 | 54.25 | 0.43307 | 48.15 | 55.35 | 1.4764 | 1.0433 | 26.5 | 48 | 55.45 | 1.4173 | 0.98425 | 25
70 | 19372.2667 | 19.3723 | 15411.5089 | 45.3 | 56 | 0.51181 |  | 47.15 | 57.8 | 1.9882 | 1.4764 |  | 37.5 |  | 49 | 59.65 | 3.4449 | 2.9331 |  | 74.5 |  | 50.4 | 55.8 | 2.5 | 1.9882 | 50.5 | 48.5 | 59.1 | 3.0315 | 2.5197 |  | 64 | 49.75 | 55.2 | 2.0079 | 1.4961 | 38 | 47.9 | 58.5 | 2.5492 | 2.0374 |  | 51.75 | 50.05 | 55.5 | 2.2244 | 1.7126 | 43.5 | 48.35 | 58.9 | 2.8642 | 2.3524 |  | 59.75 | 49 | 59.5 | 3.3465 | 2.8346 |  | 72 | 49.2 | 57.85 | 2.8543 | 2.3425 | 59.5 | 49.8 | 55.25 | 2.0276 | 1.5157 | 38.5 | 48.75 | 57.45 | 2.5 | 1.9882 | 50.5 | 49 | 57.7 | 2.6969 | 2.185 | 55.5 | 52.7 | 61.4 | 5.5906 | 5.0787 | 129 | 59.2 | 67.6 | 10.5906 | 10.0787 | 256 | 50.55 | 59.2 | 3.878 | 3.3661 | 85.5 | 46.8 | 57.7 | 1.7913 | 1.2795 | 32.5 | 45.55 | 56.5 | 0.82677 | 0.31496 | 8 | 46.5 | 57.45 | 1.5945 | 1.0827 | 27.5 | 45.95 | 56.85 | 1.1417 | 0.62992 | 16 | 46.9 | 57.75 | 1.8701 | 1.3583 | 34.5 | 46.45 | 57.3 | 1.5157 | 1.0039 | 25.5 | 48.2 | 59 | 2.874 | 2.3622 | 60 | 45.75 | 56.6 | 0.96457 | 0.45276 | 11.5 | 46.85 | 54.35 | 0.5315 | 48.35 | 55.55 | 1.6339 | 1.1024 | 28 | 48.25 | 55.7 | 1.6142 | 1.0827 | 27.5
75 | 20756 | 20.756 | 16512.3309 | 45.4 | 56.1 | 0.59055 |  | 47.3 | 58 | 2.126 | 1.5354 |  | 39 |  | 49.35 | 60.1 | 3.7598 | 3.1693 |  | 80.5 |  | 50.65 | 56.05 | 2.6969 | 2.1063 | 53.5 | 48.9 | 59.45 | 3.3268 | 2.7362 |  | 69.5 | 50 | 55.45 | 2.2047 | 1.6142 | 41 | 48.2 | 58.7 | 2.7461 | 2.1555 |  | 54.75 | 50.3 | 55.8 | 2.4409 | 1.8504 | 47 | 48.75 | 59.25 | 3.1594 | 2.5689 |  | 65.25 | 49.3 | 59.85 | 3.6024 | 3.0118 |  | 76.5 | 49.45 | 58.2 | 3.0906 | 2.5 | 63.5 | 50 | 55.45 | 2.185 | 1.5945 | 40.5 | 49.1 | 57.85 | 2.7953 | 2.2047 | 56 | 49.25 | 58 | 2.9134 | 2.3228 | 59 | 53.35 | 62 | 6.0827 | 5.4921 | 139.5 | [5 empty] | 51 | 59.65 | 4.2323 | 3.6417 | 92.5 | 47 | 57.95 | 1.9685 | 1.378 | 35 | 45.7 | 56.65 | 0.94488 | 0.35433 | 9 | 46.75 | 57.65 | 1.7717 | 1.1811 | 30 | 46.2 | 57.1 | 1.3386 | 0.74803 | 19 | 47.05 | 57.95 | 2.0079 | 1.4173 | 36 | 46.7 | 57.55 | 1.7126 | 1.122 | 28.5 | 48.55 | 59.35 | 3.1496 | 2.5591 | 65 | 45.95 | 56.75 | 1.1024 | 0.51181 | 13 | 46.95 | 54.45 | 0.61024 | 48.6 | 55.8 | 1.8307 | 1.2205 | 31 | 48.45 | 55.95 | 1.7913 | 1.1811 | 30
80 | 22139.7333 | 22.1397 | 17613.153 | 45.5 | 56.2 | 0.66929 |  | 47.6 | 58.3 | 2.3622 | 1.6929 |  | 43 |  | 49.8 | 60.45 | 4.0748 | 3.4055 |  | 86.5 |  | 51 | 56.35 | 2.9528 | 2.2835 | 58 | 49.2 | 59.85 | 3.6024 | 2.9331 |  | 74.5 | 50.3 | 55.7 | 2.4213 | 1.752 | 44.5 | 48.45 | 59 | 2.9626 | 2.2933 |  | 58.25 | 50.6 | 56.05 | 2.6575 | 1.9882 | 50.5 | 49.1 | 59.65 | 3.4547 | 2.7854 |  | 70.75 | 49.7 | 60.25 | 3.9173 | 3.248 |  | 82.5 | 49.75 | 58.5 | 3.3268 | 2.6575 | 67.5 | 50.35 | 55.75 | 2.4409 | 1.7717 | 45 | 49.4 | 58.15 | 3.0315 | 2.3622 | 60 | 49.55 | 58.3 | 3.1496 | 2.4803 | 63 | 53.95 | 62.65 | 6.5748 | 5.9055 | 150 | [5 empty] | 51.4 | 60.1 | 4.5669 | 3.8976 | 99 | 47.3 | 58.2 | 2.185 | 1.5157 | 38.5 | 45.9 | 56.8 | 1.0827 | 0.41339 | 10.5 | 47 | 57.9 | 1.9685 | 1.2992 | 33 | 46.4 | 57.25 | 1.4764 | 0.80709 | 20.5 | 47.3 | 58.25 | 2.2244 | 1.5551 | 39.5 | 46.9 | 57.75 | 1.8701 | 1.2008 | 30.5 | 48.9 | 59.7 | 3.4252 | 2.7559 | 70 | 46.1 | 56.95 | 1.2402 | 0.57087 | 14.5 | 47.05 | 54.6 | 0.70866 | 48.85 | 56 | 2.0079 | 1.2992 | 33 | 48.7 | 56.15 | 1.9685 | 1.2598 | 32
85 | 23523.4667 | 23.5235 | 18713.9751 | 45.6 | 56.35 | 0.76772 |  | 47.9 | 58.6 | 2.5984 | 1.8307 |  | 46.5 |  | 50.2 | 60.85 | 4.3898 | 3.622 |  | 92 |  | 51.3 | 56.7 | 3.2087 | 2.4409 | 62 | 49.7 | 60.25 | 3.9567 | 3.189 |  | 81 | 50.55 | 56 | 2.6378 | 1.8701 | 47.5 | 48.8 | 59.35 | 3.2382 | 2.4705 |  | 62.75 | 50.9 | 56.35 | 2.8937 | 2.126 | 54 | 49.4 | 60 | 3.7106 | 2.9429 |  | 74.75 | 50.1 | 60.65 | 4.2323 | 3.4646 |  | 88 | 50.2 | 58.9 | 3.6614 | 2.8937 | 73.5 | 50.6 | 56.05 | 2.6575 | 1.8898 | 48 | 49.7 | 58.4 | 3.248 | 2.4803 | 63 | 49.9 | 58.65 | 3.4252 | 2.6575 | 67.5 | 54.55 | 63.2 | 7.0276 | 6.2598 | 159 | [5 empty] | 51.9 | 60.55 | 4.9409 | 4.1732 | 106 | 47.5 | 58.45 | 2.3622 | 1.5945 | 40.5 | 46 | 56.95 | 1.1811 | 0.41339 | 10.5 | 47.2 | 58.15 | 2.1457 | 1.378 | 35 | 46.6 | 57.45 | 1.6339 | 0.86614 | 22 | 47.55 | 58.5 | 2.4213 | 1.6535 | 42 | 47.05 | 57.95 | 2.0079 | 1.2402 | 31.5 | 49.2 | 60.05 | 3.6811 | 2.9134 | 74 | 46.25 | 57.1 | 1.3583 | 0.59055 | 15 | 47.15 | 54.65 | 0.76772 | 49 | 56.2 | 2.1457 | 1.378 | 35 | 48.9 | 56.35 | 2.126 | 1.3583 | 34.5
90 | 24907.2 | 24.9072 | 19814.7971 | 45.8 | 56.5 | 0.90551 |  | 48.1 | 58.85 | 2.7756 | 1.8701 |  | 47.5 |  | 50.6 | 61.25 | 4.7047 | 3.7992 |  | 96.5 |  | 51.6 | 56.95 | 3.4252 | 2.5197 | 64 | 49.95 | 60.55 | 4.1732 | 3.2677 |  | 83 | 50.8 | 56.25 | 2.8346 | 1.9291 | 49 | 49 | 59.6 | 3.4154 | 2.5098 |  | 63.75 | 51.2 | 56.55 | 3.0906 | 2.185 | 55.5 | 49.8 | 60.5 | 4.065 | 3.1594 |  | 80.25 | 50.55 | 61.05 | 4.5669 | 3.6614 |  | 93 | 50.55 | 59.25 | 3.937 | 3.0315 | 77 | 50.9 | 56.3 | 2.874 | 1.9685 | 50 | 50 | 58.75 | 3.5039 | 2.5984 | 66 | 50.25 | 59 | 3.7008 | 2.7953 | 71 | 55.1 | 63.7 | 7.4409 | 6.5354 | 166 | [5 empty] | 52.3 | 61.1 | 5.315 | 4.4094 | 112 | 47.75 | 58.7 | 2.5591 | 1.6535 | 42 | 46.2 | 57.1 | 1.3189 | 0.41339 | 10.5 | 47.4 | 58.35 | 2.3031 | 1.3976 | 35.5 | 46.75 | 57.7 | 1.7913 | 0.88583 | 22.5 | 47.8 | 58.7 | 2.5984 | 1.6929 | 43 | 47.35 | 58.2 | 2.2244 | 1.3189 | 33.5 | 49.6 | 60.4 | 3.9764 | 3.0709 | 78 | 46.45 | 57.3 | 1.5157 | 0.61024 | 15.5 | 47.25 | 54.8 | 0.86614 | 49.3 | 56.45 | 2.3622 | 1.4961 | 38 | 49.15 | 56.6 | 2.3228 | 1.4567 | 37
95 | 26290.9333 | 26.2909 | 20915.6192 | 45.9 | 56.65 | 1.0039 |  | 48.45 | 59.15 | 3.0315 | 2.0276 |  | 51.5 |  | 51 | 61.65 | 5.0197 | 4.0157 |  | 102 |  | 51.9 | 57.3 | 3.6811 | 2.6772 | 68 | 50.45 | 61.1 | 4.5866 | 3.5827 |  | 91 | 51.2 | 56.7 | 3.1693 | 2.1654 | 55 | 49.4 | 60 | 3.7303 | 2.7264 |  | 69.25 | 51.5 | 56.9 | 3.3465 | 2.3425 | 59.5 | 50.15 | 60.7 | 4.2815 | 3.2776 |  | 83.25 | 50.95 | 61.45 | 4.8819 | 3.878 |  | 98.5 | 50.9 | 59.6 | 4.2126 | 3.2087 | 81.5 | 51.15 | 56.55 | 3.0709 | 2.0669 | 52.5 | 50.4 | 59.2 | 3.8386 | 2.8346 | 72 | 50.5 | 59.4 | 3.9567 | 2.9528 | 75 | 55.8 | 64.4 | 7.9921 | 6.9882 | 177.5 | [5 empty] | 52.85 | 61.65 | 5.748 | 4.7441 | 120.5 | 47.95 | 58.95 | 2.7362 | 1.7323 | 44 | 46.4 | 57.3 | 1.4764 | 0.47244 | 12 | 47.65 | 58.55 | 2.4803 | 1.4764 | 37.5 | 46.95 | 57.85 | 1.9291 | 0.9252 | 23.5 | 48.05 | 59 | 2.815 | 1.811 | 46 | 47.5 | 58.4 | 2.3622 | 1.3583 | 34.5 | 50 | 60.8 | 4.2913 | 3.2874 | 83.5 | 46.6 | 57.45 | 1.6339 | 0.62992 | 16 | 47.45 | 54.95 | 1.0039 | 49.5 | 56.65 | 2.5197 | 1.5157 | 38.5 | 49.4 | 56.8 | 2.5 | 1.4961 | 38
100 | 27674.6667 | 27.6747 | 22016.4413 | 46.1 | 56.8 | 1.1417 |  | 48.75 | 59.4 | 3.248 | 2.1063 |  | 53.5 |  | 51.4 | 62.05 | 5.3346 | 4.1929 |  | 106.5 | [6 empty] | 50.7 | 61.3 | 4.7638 | 3.622 |  | 92 | [5 empty] | 49.7 | 60.3 | 3.9665 | 2.8248 |  | 71.75 | [11 empty] | 51.35 | 61.8 | 5.1772 | 4.0354 |  | 102.5 | 51.15 | 59.85 | 4.4094 | 3.2677 | 83 | [30 empty] | 48.2 | 59.15 | 2.9134 | 1.7717 | 45 | 46.6 | 57.5 | 1.6339 | 0.49213 | 12.5 | 47.9 | 58.8 | 2.6772 | 1.5354 | 39 | 47.1 | 58.1 | 2.0866 | 0.94488 | 24 | 48.3 | 59.2 | 2.9921 | 1.8504 | 47 | 47.8 | 58.65 | 2.5787 | 1.437 | 36.5 | 50.3 | 61.15 | 4.5472 | 3.4055 | 86.5 | 46.85 | 57.65 | 1.811 | 0.66929 | 17 | 47.55 | 56.05 | 1.4764 | 49.75 | 56.95 | 2.7362 | 1.2598 | 32 | 49.65 | 57.05 | 2.6969 | 1.2205 | 31
Sample diameter: 40 mm | [5 empty] | pink | shows flow rates below the guaranteed accuracy of the flow meter | [13 empty] | More notes: | [114 empty] | comment: nominally identical to last time |  |  | Comment: slightly higher than before, but still ASTM 2/3 |  |  |  |  | Comment: exactly the same pressure differential at 20.5 L/min as previous test.
Area: | 0.0012566 | m2 | [18 empty] | Measurements taken at temperatures significantly below standard 21 oC and 85% RH. | [114 empty] | leak test performed
[6 empty] | Note: ASTM F2100 refers to mm H2O differential pressure in a test apparatus with 4.9 cm2 area and 8 L/min. | [14 empty] | No attempt to adjust volumetric flow rates for pressure change have been made, but measurement of flow on both high and low sides of the sample are in agreement to within accuracy of the gauges.
[6 empty] | That corresponds to a SLPM/M2 value of |  |  |  |  | 1.6327 | SLPM/cm2 | [8 empty] | While standard methods more often call for testing under negative pressure, these results are conducted under positive pressure- a small number of tests under negative pressure have not shown a noteworthy difference.
[11 empty] | 16326.5306 | SLPM/m2
ASTM F2100 says levels 1,2 and 3 masks must be under 5, 6 mm/cm^2, for a 4.9 cm^2 sample. That means below 24.5 mm and 29.4 mm, respectively, for this velocity.
[6 empty] | Materials with pressure drop lower than 40 mm H2O have | [6 empty] | bold | border around those pressure drops
[6 empty] | Materials passing level 2,3 rating (29.4 mm) are highlighted in | [6 empty] | green
[6 empty] | Materials passing level 1 rating (24.5 mm) are highlighted in | [6 empty] | orange

## Raw Data Clean
Pressure Drops
L/Min | Flux | H400 2x ply | H400 4x Ply | H400 (1xply, blue only) + HVMerv14 | H400 (2x ply) + HVMerv14 | H300 (1x ply, blue only) + HVMerv14 | H300 (2x ply) +HV Merv 14 | H600 (1x ply, blue only) + HVMerv14 | H600 (2x ply) + HVMerv 14 | H600, (4x ply) | HVMerv 14 (2x ply) | H100 (1x ply, blue only) + HVMerv14 | H100 (2x ply) + HVMerv 14 | H100, (4x ply) | Camfil HEPA filter | Camfil HEPA (2x ply) | Woodbridge (5x ply) | Q1 | Q2 | Q3 (retest) | Q4 | Q5 | Yellow pad, 2-layer (retest) | Yellow pad, 4-layer (4 x yellow-white back to back) | Yellow pad, 1-layer (yellow-white flow from yellow to white)
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1.3837 | 1100.8221 | 0 | 1 | 0.5 | 1 | 0 | 0.25 | -0.5 | 0.25 | 0.5 | 0 | 0 | 0.5 | 0 | 3.5 | 7 | 1 | 0 | 0 | 0 | 0 | 0 | 1.5 | 0.5 | 0
2.7675 | 2201.6441 | 3.5 | 6.5 | 3 | 5.5 | 3 | 4.25 | 2.5 | 4.75 | 6 | 5 | 2.5 | 4 | 4 | 13 | 24.5 | 7 | 2 | 0 | 1.5 | 0 | 2.5 | 2 | 5.5 | 1
4.1512 | 3302.4662 | 6 | 12.5 | 7 | 10.5 | 6 | 7.75 | 6 | 8.75 | 11 | 10.5 | 6 | 10 | 9 | 22.5 | 43 | 13.5 | 5 | 1 | 4.5 | 2 | 5.5 | 5 | 9.5 | 1.5
5.5349 | 4403.2883 | 9 | 17 | 11.5 | 14.5 | 9 | 11.75 | 9.5 | 13.25 | 16 | 13.5 | 8.5 | 12.5 | 13.5 | 31 | 61.5 | 20.5 | 7.5 | 2 | 6.5 | 4.5 | 8.5 | 7 | 14.5 | 3.5
6.9187 | 5504.1103 | 12 | 22.5 | 15.5 | 19 | 13.5 | 16.75 | 13 | 17.75 | 21 | 18 | 11.5 | 15.5 | 17.5 | 40.5 | 79.5 | 26 | 10 | 3 | 9.5 | 5 | 10.5 | 10.5 | 18.5 | 4.5
8.3024 | 6604.9324 | 17 | 27 | 19 | 23.5 | 14.5 | 18.25 | 16.5 | 22.75 | 26.5 | 23 | 14.5 | 19.5 | 21 | 50.5 | 97 | 33.5 | 12.5 | 3.5 | 11 | 6.5 | 13.5 | 12 | 22.5 | 5.5
9.6861 | 7705.7544 | 17.5 | 33.5 | 22 | 28 | 17.5 | 21.75 | 19 | 26.25 | 31 | 26.5 | 17 | 23.5 | 25.5 | 58.5 | 114.5 | 38.5 | 15.5 | 4 | 12.5 | 8 | 16 | 14 | 27 | 6
11.0699 | 8806.5765 | 19 | 38.5 | 26.5 | 33 | 20.5 | 25.25 | 23.5 | 31.25 | 37 | 31 | 20.5 | 27 | 29.5 | 67 | 137 | 45 | 17 | 5 | 14.5 | 9.5 | 18 | 15 | 32 | 7.5
12.4536 | 9907.3986 | 22 | 44.5 | 29.5 | 37.5 | 23.5 | 29.25 | 26 | 35.25 | 42 | 35.5 | 23 | 30 | 33 | 76.5 | 152.5 | 51 | 19.5 | 5.5 | 16.5 | 10.5 | 20.5 | 18 | 35.5 | 8
13.8373 | 11008.2206 | 24 | 48.5 | 32.5 | 42 | 25 | 31.75 | 28.5 | 39.25 | 47.5 | 38 | 25.5 | 33 | 37 | 85 | 172.5 | 56 | 20.5 | 6 | 18.5 | 11 | 22.5 | 19 | 39.5 | 8
15.2211 | 12109.0427 | 27.5 | 55 | 39.5 | 48.5 | 28.5 | 37.25 | 33 | 43.75 | 54 | 45 | 29 | 39.5 | 42 | 97.5 | 188.5 | 63.5 | 23.5 | 6.5 | 20.5 | 12.5 | 25.5 | 22 | 45 | 9.5
16.6048 | 13209.8648 | 31 | 62.5 | 41.5 | 54 | 32.5 | 41.25 | 36.5 | 50.75 | 60.5 | 49.5 | 32 | 43.5 | 46.5 | 108.5 | 214 | 72 | 27.5 | 8 | 23 | 14.5 | 28.5 | 24 | 50.5 | 10.5
17.9885 | 14310.6868 | 34 | 68 | 45.5 | 59 | 34.5 | 45.75 | 39.5 | 55.25 | 65 | 55 | 34.5 | 48 | 50.5 | 119 | 236 | 78.5 | 29.5 | 8 | 25 | 15.5 | 31 | 26.5 | 55.5 | 11
19.3723 | 15411.5089 | 37.5 | 74.5 | 50.5 | 64 | 38 | 51.75 | 43.5 | 59.75 | 72 | 59.5 | 38.5 | 50.5 | 55.5 | 129 | 256 | 85.5 | 32.5 | 8 | 27.5 | 16 | 34.5 | 28 | 60 | 11.5
20.756 | 16512.3309 | 39 | 80.5 | 53.5 | 69.5 | 41 | 54.75 | 47 | 65.25 | 76.5 | 63.5 | 40.5 | 56 | 59 | 139.5 |  | 92.5 | 35 | 9 | 30 | 19 | 36 | 31 | 65 | 13
22.1397 | 17613.153 | 43 | 86.5 | 58 | 74.5 | 44.5 | 58.25 | 50.5 | 70.75 | 82.5 | 67.5 | 45 | 60 | 63 | 150 |  | 99 | 38.5 | 10.5 | 32 | 20.5 | 39.5 | 33 | 70 | 14.5
23.5235 | 18713.9751 | 46.5 | 92 | 62 | 81 | 47.5 | 62.75 | 54 | 74.75 | 88 | 73.5 | 48 | 63 | 67.5 | 159 |  | 106 | 40.5 | 10.5 | 34.5 | 22 | 42 | 35 | 74 | 15
24.9072 | 19814.7971 | 47.5 | 96.5 | 64 | 83 | 49 | 63.75 | 55.5 | 80.25 | 93 | 77 | 50 | 66 | 71 | 166 |  | 112 | 42 | 10.5 | 37 | 22.5 | 43 | 38 | 78 | 15.5
26.2909 | 20915.6192 | 51.5 | 102 | 68 | 91 | 55 | 69.25 | 59.5 | 83.25 | 98.5 | 81.5 | 52.5 | 72 | 75 | 177.5 |  | 120.5 | 44 | 12 | 38 | 23.5 | 46 | 38.5 | 83.5 | 16
27.6747 | 22016.4413 | 53.5 | 106.5 |  | 92 |  | 71.75 |  |  | 102.5 | 83 | [6 empty] | 45 | 12.5 | 31 | 24 | 47 | 32 | 86.5 | 17
Sample diameter: 40 mm
Area: | 0.0012566 | m2
Note: ASTM F2100 refers to mm H2O differential pressure in a test apparatus with 4.9 cm2 area and 8 L/min.
|  |  |  | That corresponds to a SLPM/M2 value of | 8 | 1.6327 | SLPM/cm2
[5 empty] | 8 | 16326.5306 | SLPM/m2
[5 empty] | 4 | 0.81633 | SLPM/cm2
[5 empty] | 4 | 8163.2653 | SLPM/m2
ASTM F2100 says levels 1,2 and 3 masks must be under 5, 6 mm/cm^2, for a 4.9 cm^2 sample. That means below 24.5 mm and 29.4 mm, respectively, for this velocity.
|  |  |  | Level 1 | 24.5 mm | Green
|  |  |  | Level 2,3 | 29.5 mm | Orange

## Relevant Flux Runs
Flux
|  | 15411.5089 | 16512.3309 | 7705.7544 | 8806.5765
|  | 3-4 L/min | 3-4 L/min | 7-8 L/min | 7-8 L/min
Pressure Drop | Q2 | 4 | 5 | 8 | 9
| Yellow pad, 1-layer (yellow-white flow from yellow to white) | 6 | 7.5 | 11.5 | 13
| Q4 | 8 | 9.5 | 16 | 19
| Q3 (retest) | 12.5 | 14.5 | 27.5 | 30
| Yellow pad, 2-layer (retest) | 14 | 15 | 28 | 31
| Q1 | 15.5 | 17 | 32.5 | 35
| Q5 | 16 | 18 | 34.5 | 36
| H100 (1x ply, blue only) + HVMerv14 | 17 | 20.5 | 38.5 | 40.5
| H400 2x ply | 17.5 | 19 | 37.5 | 39
| H300 (1x ply, blue only) + HVMerv14 | 17.5 | 20.5 | 38 | 41
| H600 (1x ply, blue only) + HVMerv14 | 19 | 23.5 | 43.5 | 47
| H300 (2x ply) +HV Merv 14 | 21.75 | 25.25 | 51.75 | 54.75
| H400 (1xply, blue only) + HVMerv14 | 22 | 26.5 | 50.5 | 53.5
| H100 (2x ply) + HVMerv 14 | 23.5 | 27 | 50.5 | 56
| H100, (4x ply) | 25.5 | 29.5 | 55.5 | 59
| H600 (2x ply) + HVMerv 14 | 26.25 | 31.25 | 59.75 | 65.25
| HVMerv 14 (2x ply) | 26.5 | 31 | 59.5 | 63.5
| Yellow pad, 4-layer (4 x yellow-white back to back) | 27 | 32 | 60 | 65
| H400 (2x ply) + HVMerv14 | 28 | 33 | 64 | 69.5
| H600, (4x ply) | 31 | 37 | 72 | 76.5
| H400 4x Ply | 33.5 | 38.5 | 74.5 | 80.5
| Woodbridge (5x ply) | 38.5 | 45 | 85.5 | 92.5
| Camfil HEPA filter | 58.5 | 67 | 129 | 139.5
| Camfil HEPA (2x ply) | 114.5 | 137 | 256